## Sheet1
TBI
inst nr | Question 7 : For which indications is TBI used in your institute? | Question 8 : Do you have a dedicated TBI team? | Question 9 : Do you have a backup treatment machine | Question 10 : Do you have to allocate a time slot for treatment? | Question 11 : How many patients do you treat annually with TBI? | Question 12 : Planning technique | Question 13 : CT-Planning: please describe patient positioning and start/end point of CT, slice thickness | Question 14 : Number of beams | Question 15 : Any additonal informantion on the number of beams | Question 16 : Is a Field-in-field technique used? (supplemental small fields to homogenise the dose) | Comments | Question 17 : What beam/patient orientation is used? (eg. gantry at 90 degrees and the patients lying at their side, therefore AP/PA for the patient). Please be elaborate. | Question 18 : What beam quality/energy is used (MV) | Question 19 : What are the dose rates used (cGy/min midplane). Please specify per dose level (eg. 1x8 Gy or 2x 6Gy). | Question 20 : What is the patient position (e.g. standing, lying (prone, supine, left side, right side), dedicated chair..) | Question 21 : Do you use a fixed focus-skin or a fixed focus-midline distance technique? | Question 22 : Focus-Skin/Midline distance for treatment [cm] | Question 23 : Source-Axis-Distance (SAD) treament unit [cm] | Question 24 : Field size at iso [cm] | Question 25 : Field orientation (collimator angle) | Question 26 : Describe your pre-treatment QA: what tolerances are allowed and what actions are performed. | Question 27 : Is in vivo dosimetry performed on a routine basis? | Question 28 : What detectors are used for in vivo dosimetry? | Question 29 : What positions are measured in vivo? | Question 30 : What are the tolerances on the in vivo measurements? | Question 31 : Do you use In-Vivo dosimetry non-routinely (eg. while introducing a new technique)? Please elaborate. | Question 32 : What calculation algorithm and inhomogeneity correction are used? | Question 33 : Is a manual (table based) or CT-based MU calculation performed? Please elaborate. | Question 34 : Is shielding of Organs At Risk used? | Question 35 : Is individual or standardized shielding used? | Question 36 : How is the thickness of the shielding determined? | Question 37 : How are the shielding devices fixated? | Question 38 : What normalization method is used (point, dvh,..)? | Question 39 : Is dose inhomogeneity taken into account? If so, what tolerances are allowed? | Question 40 : How is the treatment data entered into the R/V system (automatic, manual, combined (both manual and automatic), standard plans) | Question 41 : Describe Patient positioning, fixation and verification, especially extremities (arms, legs, head) | Question 42 : Describe position verification of shielding and other tools/devices | Question 43 : Describe distances between machine and shielding devices | Question 44 : Do you use additional beam modifying devices (e.g. “scatter/spoiler-screen”)? Please describe sizes, materials and position. | Question 45 : What hygienic precautions do you take? | Question 46 : What personnel is available during treatment and what is their role? | Question 47 : What personnel is required for pre-treatment checks and what is their role? | Question 48 : What published guidelines did you use (AAPM, IAEA, etc..) | Question 49 : Have you used scientific publications (please state which)? | Question 50 : Have you performed a prospective risk analysis while setting up your treatment? (include e.g. lasers, cross-hairs,..) | Question 51 : Which method of risk analysis did you use? (e.g. HFMEA, ..) | Question 52 : Which prominent risks were indicated in the analysis? | Question 53 : Please add here any additional relevant information.
1 | According HOVON guidelines for indications and transplant schedule: myelo-ablative regime: 6x2Gy/3 days, 6h time interval, shielding. Very rarely applied, once a year? non-myelo-ablative regime: 1 or 2x2Gy. No shielding of OAR. This scheme most often used. see appendix. | Yes | On site | Regularly scheduled | 27 | ''Manual'' (combination of pdd tables and some kind of thickness determination) |  | 2 | AP and PA | No |  | TBI-chair:gantry 270 patient in TBI-chair, chair rotates between AP and PA Table: gantry 270 patient &lt;150 cm length at table, patients lying on left side and right side | 10 | for all: ~ 25 cGy/min (per beam, not corrected for the time between beams) | TBI-chair:dedicated chair to increase homogeneity in effective depth (zweedse stoel), knees bended. Neck slightly bended forward. Table: left side and right side with knees bended and neck slightly bended forward. | focus-midline | 450 | 100 | 40x40 | 45° | per linac commissioning (treatment planning) at TBI distance. Further: standard QA of system. | No |  |  |  | Long ago, at introduction of TBI-chair and treatment planning procedure. |  | We use for calculation a TBI planning-spreadsheet: Emperic formulas: based on PDD curves, measured linac parameters, geometrical measures and dimensions of the patient. | Yes | Only for 6x2Gy schedule: lung: individualized shielding kindney: individualized shielding eyes: individualized time of standard shielding | A chair: lung/kidney/eyes: calculation thickness based geometrical parameters/dimensions (using our TBI planning-spreadsheet) B: table: the same, however for lung, measure of lung thickness, and electron density as function of age, using table from literature. | lung and kidney: using tape on perspex-plate which is adjustable on front or back screen. eyes: RVS cilindrical standard shielding in adjustable hold | Normalization in mid of average thickness of 3 measures at mediastinum cranial, mediastinum central, mediastinum caudal. | Only for lungs, by using equivalent thickness. No tolerances | Beams are entered manually in Mosaiq + check. PDF of treatment planning spreadsheet added. | Chair: see added picture Knee support and seat adjustable: height, rotation and distance to screen. Measures are stored in data-base Hands on knees. Head slightly bended Table: bended knees, lower arm with hand at head | Lung and kidney using fosfor cassettes. Using 25 MU for each beam. After reading decision and further irradiation. | shielding on chair or table, thus close to body | no | cleaning chair/table before and after treatment with alcohol 70% Wearing masks, special clothing, gloves (certainly for children, but often for adults as well) | RTT's: 3 involved in set-up and treatment Physicist: available Radiation Oncologist: available | Check of calculation of shielding by Medical Physics Engineer or Medical Physicist. | o.a. AAPM 1986 | for shielding and OAR dose publications by van Kempen et al. | No |  |  | See wordfile Probably better to have a telephone meeting, or live meeting to see how it is working
2 | AML-ALL: 1375 cGy in 10 fractions (3x per day) • Total abdomen midline dose : 1375 cGy +/- 5% • Total lung midline dose : 900 cGy +/- 5% • Homogeneity : +/- 5 % (10% max.) • Doserate : max. 10 cGy/min (1.66 mGy/s) • Shielding: personalized lungblocks, for every fraction 1 lung is shielded (odd fractions: left lung, even fractions: right lung) + scatter screen. MD: 2 options: 4 Gy in 2 fractions • Total abdomen midline dose : 4 Gy +/- 5% • Homogeneity : +/- 5 % (10% max.) • Doserate : max. 10 cGy/min (1.66 mGy/s) • No shielding, use of scatter screen 2 Gy in 1 fraction (Single fraction • Total abdomen midline dose : 2Gy +/- 5% • Homogeneity : +/- 5% (10% max.) • Doserate : max. 10cGy/min (1.66mGy/s) • No shielding, use of scatter screen | Yes | No | Regularly scheduled | 10 | ''Manual'' (combination of pdd tables and some kind of thickness determination) |  | 1 | Odd fractions: patient is lying at the right side Even fractions: patient is lying at the left side For all fractions the same beam is used. | No | 1 A longblock is used to shield the left long in the odd fractions. A second longblock is used to shield the right long in the even fractions. | gantry at -90°, patient is lying at his/her side (midline distance to the gantry 3.7m): odd fractions: patient is lying at the right side even fractions: patient is lying at the left side AP/PA irradiation from the patient's point of view | 18 | max 10 cGy/min (usually 4 cGy/min) | Lying at the side on a dedicated table. Knees and legs bent, head on a pillow, lower arm bent around head, upper arm stretched on hip. | focus-midline | 370 | 370 | 40x40 | 135 | Calibration of the individual OmniPro InviDos detectors using a PMMA slab-phantom with a Farmer chamber in its center. The thickness and position of the slab phantom corresponds to the patient's anatomy. The output of the invivo detectors is correlated to the absolute dose measured with the farmer at the position of the abdomen, mediastinum, lung and head of the patient. For each anatomical position, the thickness of the slab phantom corresponds to the patient's anatomical thickness. For the lungs, the equivalent thickness is calculated based on a CT scan of the patient in the treatment position. | Yes | OmniPro InViDos (IBA) | head, left lung, right lung, mediastinum, abdomen | 5% |  | Not applicable | The MU calculation is based on the absolute dose measured with a farmer chamber in the center of the slab phantom. The slab phantom is positioned on the treatment table at the position of the patients abdomen and its thickness corresponds to the thickness of the patient's abdomen. | Yes | Individual lung blocks | For the lung blocks, a standard thickness is used. Depending on the lung dose, measured in vivo, the number of MU for which the block is used, is adjusted. This means that the block is removed earlier in the last fractions or an additional lung block is made and added during the last fractions. | The lung block is attached to the scatter screen which is positioned in front of the patient. | Not applicable | During the irradiation, invivo dosimetry is performed using the Invidos detectors at the position of the patients abdomen, upper and lower lung, mediastinum and head. If the measurements show a deviation &gt;5% from the prescribed dose, the total number of MU can be decreased, a lung block can be added/removed or the distance of a particular part of the body to the gantry can be altered (e.g. head closer to the gantry). | Manual entry in the R/V system | The patient is positioned at the right side (odd fractions) or at the left side (even fractions). The legs are bent (to fit the toes in the 80% isodose level), the upper arm is positioned on top of the hip and the lower arm is slightly bent around the head. | The position of the lung block is verified using an X-ray radiograph. | 317 cm between the scatter screen and the gantry | A scatterscreen is used: plexiglass (?), dimensions: 70cm x 215 x 1.5 cm³ | Desinfectation of the table, pillows and devices that come into contact with the patient. | 1 physicist: positioning patient, lung block and scatter screen, positioning invivo dosimeters, operating the linac (service mode), operating the Invidos system, monitoring dose, monitoring patient and machine 1 nurse: positioning and observing patient 1 doctor: positioning patient and lung block, evaluating radiograph | 1 physicist to perform the calibration of the invivo detectors (see previous questions). 1 nurse+doctor to perform a CT scan of the patient in the treatment position. |  | Radiotherapy and Oncology 52 (1999) 269-276 In-vivo dosimetry by diode semiconductors in combination with portal films during TBI: reporting a 5-year clinical experience.P P Mangili1 , C C Fiorino1 , A A Rosso1 , G M GM Cattaneo1 , R R Parisi1 , E E Villa1 , R R Calandrino1 1Servizio di Fisica Sanitaria, H. San Raffaele, Milan, Italy Int J Radiat Oncol Biol Phys. 2007 Feb 1;67(2):544-51. Epub 2006 Dec 4 Reduced-intensity conditioning regimen using low-dose total body irradiation before allogeneic transplant for hematologic malignancies: Experience from the European Group for Blood and Marrow Transplantation.Belkacémi Y1, Labopin M, Hennequin C, Hoffstetter S, Mungai R, Wygoda M, Lundell M, Finke J, Aktinson C, Lorchel F, Durdux C, Basara N. | No
3 | a paediatrcis and adults b 6 x 2 Gy (2 fractions/day) -&gt; child 2 x 5 Gy (2 days) 8 x 1.5 Gy (2 fractions/day) 1 x 2 or 4 Gy | No | Other site | End of the day | 9 | ''Manual'' (combination of pdd tables and some kind of thickness determination) |  | 2 | We irradiate the patient with one anterior and one posterior beam. The weighting is 50% | No |  | Gantry 90, colli 315 left and right sided for patient. The patient lays down on his/her right and left side. | 15 | 2x 5 Gy and 6x 2 Gy for pediatrics 7.96 cGy / min | lying in a bed: on his/her left and right side. | focus-skin | 500 | 100 | 40x40 | 315 | no pre-treatment QA during treatment the dose is monitored (TLDs) After each fraction we adjust the calculations to obtain a dose within the tolerance level of 5% | Yes | TLD | head, neck, lungs (upper and lower lungs), abdomen and legs | +/- 5 % | no but we want to start with diodes and we sometimes try to use new dosimeters (in house developed mosfets and verify them with the TLDs) | We correct for the path length in the lungs. | Manual, table based | Yes | individual | measurements on CT, correction for the path length trough tissue and lungs | with tape | point (middle of the abdomen) | No | manual | Patient is lying down in a bed on his/her right side (anterior beam). He/she has to lay down as flat as possible to the plastic plate on his/her back. The knees are a little bit bended. The right arm supports the head (the arm is positioned under the pillow). The left arm is laying down on the left him, this arm is in an extended position. | The shielding of the blocks (lungs and head and neck) is verified with an KV image. | 500 cm | No | a mouth mask for everyone. And the material that is used has to be sterile or cleaned with alcohol. | 1 physician, 1 physicist, 1 nurse. Physicist is responsible for the positioning of the patient and the treatment parameters. The physician is responsible for the shielding at the right place and the nurse is there to support the patient | Physician, physicist and nurse. Physician: to calculate the shape of the lung shielding Physicist: to calculate the thickness of the lung shielding and the number of MU Nurse: to support the patient |  | 0 | No |  |  | We want to update our technique, because we want to use an other linac to treat our TBI patients.
4 | a. conditional scheme for agressive DLBCL, AML, ALL, MDS, allogene stemcelltransplantation. Only one prescription: 12Gy. b. 6x2Gy/3days at midline abdomen, at least 6hours interval c. Radiotherapy start at day-8 and ends at day-5. Day 0 is the day of stemcelltransplantation. d. No deviation from time scheme are allowed. | Yes | On site | Regularly scheduled | 6 | ''Manual'' (combination of pdd tables and some kind of thickness determination) |  | 2 | Anterior-Posterior beams | No |  | Gantry at 272°, patient is position half-way between standing/sitting such that the horizontal projection of the thickness of legs and arms is comparable to the horizontal projection of the thickness of the abdomen. the patient changes of orientation to get an AP/PA irradiation | 6 | 16cGy/min at midplane for 6x 2Gy. | A dedicated chair is used to get the position half-way standing/sitting. | focus-midline | 557 | 100 | 40x40 | 45 | Lung shielding has to be clinically align to the lungs. The physician determines himself of the alignment is acceptable. | Yes | diode | at umbilicus | 10% | We use Gafchromic EBT3 stripes of film for new techniques (surface dosimetry). |  | Table based calculation based on 1/r² correction and TPR correction. | Yes | Individual shielding | to get 75% transmission (8mm of lead) | with tape on a vertical plexiplate | point at midplane | No, it is not taken into account. | Manually adapted from a standard plan | Depending of the size and thickness of the patient, the height of the seat and the height of the plate for the feet are adapted such that the horizontal projection of the thickness of legs is similar to the horizontal projection of the thickness of the abdomen. | Radiological RX-plates are position behind the thorax of the patient to control the correct alignment of lung-shielding. | 500cm between source and lung-shielding | We use a scatter/spoiler-screen of 2cm plexiglas at 500cm from the source. Lung-shielding is taped op the scatter/spoiler-screen. | Patient specific protection measures (mouth-mask) | RTTS position the patient, the lung-shieldings, position the diodes and deliver the treatment. Physician and physicist are available upon request. | RTTS position the patient, the lung-shieldings, the radiological RX-plates and provide the imaging results. The physician controls the alignment of lung-shielding. | AAPM report 17 | / | No |  |  | /
5 | Malignant disease: •leukemea (ALL, ANLL, CML) •myelodysplastic syndrome (MDS) •lymphoma •solid tumors •multiple myeloma Benign disease: •severe combined immune deficiency (SCID) •aplastic anaemea •Fanconi anaemea •severe progressive systemic and polyarticular juvenile chronic arthritis •Wiskott-Aldrich •Thalassemea | Yes | On site | Regularly scheduled | 25 | CT Based | 2 CT's in left and right side position are used for lung block shaping and thickness; also to determine the patient on different levels. Scans 5 mm coupes chin till coccyx. PDD tables are used for calculating MUs and doses at different levels | 2 | AP and PA | Yes | 1 if dose inhomogeneity is too high (obese patients) we add field in field to homogenise dose. Perhaps in the future full MLC planning Maastricht/Nijmegen 'style' | gantry 90 (or 270) patient lying on left and right side during treatment, thus AP/PA for patient | 6 | we always use max smoke, say 550 MU/min. A central TBI dose of 9 Gy typically corresponds to 2 x 8500 MU (17.000 MU). Hence the doserate is typically 0.3 Gy/min | half treatment left and half treatment right side | focus-midline | 400 | 100 | 40 cm x 40 cm | 45 degree | no pre treatment QA, only standard linac QA | No |  |  |  | no | NCS 12 plus equivalent path length for lung | manual based MU calculation based on PDD measurements at 4 m | Yes | individual lungblock and standardized eye shielding | we determine the lungdose without blocks and from that point we calculate the required lung block thickness to reduce the lung dose to 6 gy | lung blocks at 2 m (the lung blocks are mounted onto a perspex plate that is positioned on the treatment table with a plateau rotation to enable positioning at 2 m). Eye shielding in a standard just in front of patient (midplane patient at 4 m, hence eyeshielding at approximately 3.9 m) | normalization to the average point dose at sternum, belly button and hip height | if the central doses are outside the range 90%-115% field in field is applied to reduce dose inhomogeneity | via Pinnacle (MU and plan from tables setup on waterphantom in Pinnacle, this plan is sent to Mosaiq) | patient lying on left and right side upper arm along the body and lower arm supporting head, in a bowed manner | Cablon TBI imager is used for verification of the lung shielding. The position of the block versus vertebra and lung is checked on the imager, both pre-treatment (with a small fraction of the field, 20 MU), and during treatment (real-time imaging). | lungblocks are at 2 m from focus. Eye shielding just in front of the patient. Since the midplane patient is at 4 m the eyeshielding is at approximately 3.9 m | no | if required sterile clothing, otherwise no special precautions | radiation therapists perform setup and irradiation radiation oncologist (approval treatment setup, lung and eye shielding positioning, and check position lung shielding on radiochromic film) | radiation therapists check if every required item is available. We do not perform dosimetric pretreatment checks | AAPM 17 | PhD thesis Loes van Kempen, Optimization of total body irradiation: the match between (maximal) leukemic cell kill and (minimal) late effects | Yes | SAFER | 1 poor quality position verification lung shielding using radiochromic film. 2. intra fraction movement of patient (''sagging patient'')
6 | INDICATIE: CURATIEF Acute Myeloïde Leukemie (AML) Acute Lymfatische Leukemie (ALL) Recidief AML in CR ALL in CR Myelo Dysplastisch Syndroom Chronische Myeloïde Leukemie M. Kahler Bij een allogene stamcel transplantatie krijgt men stamcellen terug van een donor (Dit kan een sibling zijn maar ook een MUD = Matched unrelated donor) Bij een autologe stamcel transplantatie krijgt men het eigen beenmerg terug, dit wordt toegepast bij afwezigheid van een donor. DOSISSCHEMA: Non-myeloablatief: 1 x 2 Gy Hierbij wordt afgeweken als: Cord blood transplantatie: 2 x 2 Gy op 2 achtereenvolgende dagen Flamsa schema:2 x 2 Gy op 1 dag met 6 uur ertussen. Myelo-ablatief (bv bij jongere patienten): 6 x 1.67 Gy 2 maal per dag op gehele lichaam (ogen van AP afgeblokt) | Yes | On site | Regularly scheduled | 37 | ''Manual'' (combination of pdd tables and some kind of thickness determination) |  | 2 |  | No |  | Gantry 90, collimator 45 degrees. Patients on their side. AP, turn bed 180 deg, then PA. Focus-skin distance 400 cm. | 10 | 0 | lying on most comfortable side (left or right) | focus-skin | 400 | 100 | 40x40 | 45 | Ionization chamber dosimetry in a waterphantom | No |  |  |  | No | Output factor * PDD PDD depth based on measurement of diameter at different body locations. Inhomogeneity only for lungs. | Table has been made part of our self written computer (chart)program. | Yes | Standardized eye blocks in AP field | On basis of attenuation coëfficiënt of cerrobend | Fixated on a perspex plate, attached to a vertical standard | dosis at midline | no | manual | Patient on side on steady mattress, knees pulled upwards, arms alongside body. Verification with laser. No head fixation. | 0 | focus shield distance about 370 cm | Perspex spoiler screen, attached to patient bed. Size covers compleet length of patient. Some 15 cm distance from patients' foreside. |  | Radiation therapists (laboranten), radiation oncologist | Radiation therapists (laboranten) | None | None | No
7 | 1x2 Gy, autogene bonemarrow transplantation; 2x2 Gy, subsequent days, transplanation with navel cord stem cells; 2x2 Gy, twice per day 8hr interval, acute leukemia; | No | On site | No | 50 | ''Manual'' (combination of pdd tables and some kind of thickness determination) |  | 2 | 15 MV, doserate 300 MU/min, collimator 45 deg | No |  | AP-PA, patient lying on same side on dedicated bed with polystyrene ''scatter screen'', bed rotated in between | 15 | 300 MU/min; typically about 1900 MU given per beam. | Lying on side (same side for both beams, bed is rotated in between) | focus-skin | 415 | 100 | 39.9 | 45 | Determine patient diameter prior to treatment (must be within 2 cm before new MU calculation is needed); Field parameters are checked (standardized) | No |  |  |  | Yes but not for TBI: the technique used has not been altered for long time. | Manual calculation (waterbased) | Table-based | No |  |  |  | midline belly normalization | No | Partial manually (combined). Pre-calculated MUs are entered depending on diameter core patient | Patients lay on right side, left arm on side, right arm bended under head, knees somewhat bended to stabilize position | n.a. | n.a. | Spoiler-screen, 1cm thick, transparant polystyrene | Patients lay on a hospital bed of which the covers are replaced after treatment | Technicians/therapists who perform treatment. Physicist-on-call if needed (re-calculation). | Therapist/technician: verification of filled in parameters; Physicist: verification of filled in parameters | Unknown | Unknown | No |  |  | We might revise / update our treatment technique
8 | Indicaties hoog risico hematologische ziektes Wb mini TBI mn AML en MDS Wb grote TBI jongere pt &lt;40 bijv recidief Hodgkin/paragranuloom, hoog risico agressief NHL, ALL... 1-2x 2 gy indien niet myeloablatief afh soort donorcellen cq indicatie op dag -1 en evt -2 ( dag O =sct) als dat echt belangrijk moeten we protocol erbij halen Dag -1 en -2 bij 2x 6gy 6gy | Yes | On site | Regularly scheduled | 70 | ''Manual'' (combination of pdd tables and some kind of thickness determination) |  | 2 | A beam is limited to 1000 MU, so each beam (AP and PA) of approx. 5000-6000 MU are in fact 5 - 6 beams. That is, the total number of beams is 10 - 12. | No |  | Gantry 90 degrees, patient at their side, lying on a bed, behind a Perspex buildup plate, therefore AP/PA for the patient. Answer to Question 18: 23 MV (1st linac), 18 MV (backup linac), and 10 MV (2nd backup linac) | 10 | Approx. 30 - 40 cGy/min for all dose levels. | left side and right side | focus-skin | 450 | 100 | Diagonal 40 x 40 cm (i.e. with very shallow penumbra) | 45 | Yearly QA: - absolute dose at standard TBI-geometry (SDD approx. 420 cm). - beam profile at standard TBI-geometry. No patient specific pre-treatment QA. | Yes | TLD | Near head and feet. | 10% | No | Manual calculation based on patient thickness. TMR table. Equivalent pathlength method. | Manual. CT is used to determine patient thickness at several positions and distance through lung. | Yes | Individual | Thickness determined by the maximum tolerance of the organ at risk. For example for 2 x 6 Gy the lungs should receive no more than 4 Gy per fraction. Skull sparing for High BMI patients at 2x6Gy | Double sided adhesive tape | Point | No. Except for shielding of lungs and kidneys, whole body receives the same unshielded field. Therefore accepting an inhomogeneous dose. | The beams are a standard plan, but the MU (and thus the number of beams) are generated by a software program based on patient thickness etc.. The plan file is imported into the R&V-system. No manual modifications of the linac parameters are required/allowed. | Patient lying on her side. Head on a pillow upper arm above the head. Legs slightly bent. No fixation. | Imaging with a phosphor plate using the first 30-50 MU of the treatment beam. Adjust shield positions for lung and kidney shielding. Second set of images to verify corrections. Wiil use Theraview TBI imager in the future | Focus - shielding approx. 4 m. | No. | Clean clothes. | 3 RTT: positioning the patient and shielding, deliver imaging/treatment beams, readout the phosphor plates, adjust shielding. 1 physician: Judge/adjust/approve shielding. 1 physicist stand by: recalculate MUs when treatment needs to continued on backup linac (due to malfunction of the first linac). | N.a.
9 | (sorry I'm still waiting for informations from the physician in charge). a. paediatrics from time to time (1 over 10 patients) b. 6x2Gy/6days c.? d.? | No | On site | Start of the day | 36 | ''Manual'' (combination of pdd tables and some kind of thickness determination) |  | 4 |  | No |  | Gantry 270° and patients lying at their side (AP/PA) and twice supine (RL/LR) | 15 | 5-6 cGy/min midplane, 2 Gy per fraction | lying (left side, right side, supine twice) | focus-midline | 403 | 100 | 40 | 45 | No patient-specific QA. TBI commissioning: flatness, dose (reference chamber), dose rate, PDD, diode calibration. No TBI-specific QA on regular basis. | Yes | 10 semiconductor detectors | head (entrance and exit) mediastinum (entrance and exit) abdomen (entrance and exit) | 5 % with respect to expected values (calculated MUs based on phantom measurements) | no | Calculation based on commissioning measurements in a water-equivalent phantom. (dose at midline, depending on phantom thickness). Water equivalent thickness used for lung (density 0.3) | Table-based MU calculation: dose midline to mediastinum and abdomen (average) depending on patient thickness. | Yes | No | Standardized head shielding (6 cm PMMA, 87 % transmission), used 50 % of the treatment, in order to deliver 2 Gy at head level. | Head shielding (6 cm PMMA) fixed on a beam spoiler screen (2 cm PMMA) positioned 50 cm ahead of treatment bed. | 2 Gy midline at mediatinum and abdomen level (on average) | Only for lung, in a simplified way (water-equivalent). | Stored beams are delivered in service mode (no R/V system for the time being but we plan to integrate TBI treatments into MOSAIQ, once we solve a dose rate problem). | Patient lying on a treatment bed, centered in the midline as much as possible. Height of the bed was adjusted using crosshair and light field. No immobilization system. In supine position, legs are attached to a longitudinal bar, for comfort, and arms are along the body. | Head shielding is positioned using light field, in order to protect head and heck. Treatment bed is fixed, via two holes drilled in the floor. | 350 cm | We use a 2-cm thick PMMA scatter screen (transmission 94 %) positioned 50 cm ahead of treatment bed (marks on the floor). This 2.5 m x 1.5 m screen is mounted on wheels and can be easily positioned. | Two nurses are in charge. They wear gloves and masks. Detectors are cleaned up after treatment. | Two nurses: positioning of the bed, the screen, the shielding, positioning of the patient, 10 diodes are sticked on patient skin. One physicist: dedicated Excel sheet for monitor units calculation based on diodes. Must be within 5 % compared to estimated values (calculation based on PDD measured during commissioning). MUs delivered in service mode, always in the range 700-900 per field. Verification of the delivered dose rate: 100 MU/min. | Two nurses: positioning of the patient (supine, left/right side) and verification of head shielding before each field (4 times), sticking of the diodes. One physicist: verification of the setup, connection of the electrometer to the computer. One physician: in case of patient sickness | AAPM TG 29 (1986) | 4-field technique was widely spread in the 90's when we started treatments. Most interesting articles we used: -physical aspects of TBI at the Middlesex Hospital, Planskoy et al, Phys.Med.Biol. 41 (1996) 2307-2343. -The dosimetry of fixed, single source hemibody and total body irradiators, Glasgow, Med. Phys. 9 (3) 1982 311-323 - Calibration of semiconductor detectors for dose assessment in TBI, Jornet et al, Radiotherapy and Oncology 38 (1996) 247-251 - In vivo dosimetry by diode semiconductors in combination with portal films during TBI, Mangili et al, Radiotherapy and Oncology 52 (1999) 269-276 | No
10 | leukemia and Non Hodgkin lymphma: through reduced intensity conditioning for allogene stem cell donation. Chemotherapy finishes 1 day before start RT (delay not acceptable). Stem cells have to be donated maximum 48 hours after RT is given. 2 schedules: 1x2 Gy for regular stem cells 2x2 Gy (1 day interval) for umbilical cord stem cells . 4x2 and 5x2 for mixed indications, among which mantle cell lymphoma | Yes | On site | Regularly scheduled | 40 | CT Based | slice thickness 1 cm, scan length 150 cm, patient positioned in vacuum mattress with knees bent. this way we can usually scan the entire patient. sometimes a part of the toes is not in the image. furthe positioning details further on. | 4 or 6 | 2 large open fields and two mlc fields for shielding extremities, for 5x2: 6 beams 2 top-up fields for ribs | Yes | 1 see 15 | gantry at 90 and 270, patient is moved to other side of bunker both beams with head to the gantry patient lying on their side | 10 | maximum available: 600 (mu)cG/min at 1m, approximately 600/4^2~ 38 (mu)cGy/min | lying on their left side fixed in vacuum mattress, knees bent with positioning lines from ct scanning as a reference. Right arm downwards on the right side and left arm across the chest holding right shoulder. Dedicated bed adjustable in height. | focus-midline | 390 | 100 | 40 | 45 degrees | an automated mu check is performed, tolerances +/- 5%. no separate output measurement, we rely on the regular output measurements on 'normal' ssd's (after 2 years of comparing output measurements on normal and large ssd's). Annual check of output at large ssd is performed. | Yes | mosfets | neck, lung, bellybutton, lower leg. midline dose is recorded in tps, midline dose is extrapolated from entrance and exit dose measurement with mosfet. 1 mosfet on spoiler screen for linac output check, actual position screen is recorded for inverse square law correction. Data are analysed after completion of treatment, we are thinking about a possibility to have an online check but that does not seem to be practically possible. | +/- 10 %. | no other method than our standard. | we use AAA. | full 3d dose calculation in eclipse. we are currently implementing and verifying the spoiler screen in the plan to get a more accurate skin dose (has some influence on the PTV normalisation we use) | yes | yes for 5x2. With MLC | TPS (field in field block with MLC) | Gantry/MLC | we normalise on PTV coverage 95% has to cover 99% of the ptv. PTV is the entire body. | body max &lt; 150%, lung max &lt;120%. | transferred as a regular eclipse plan within the aria suite. | Patent lies on their side in a body-length vacuum mattress, with knees bent. long lines on the right side and anterior and posterior for aligning the body with the lasers. Additional ceiling mounted lasers at large SSD. Right arm downwards on the right side and left arm across the chest holding right shoulder or on top of head (5x2 Gy). Dedicated bed adjustable in height. | No shielding used (only 2 Gy or 2x2Gy). Spoiler screen as close to the patient as possible and parallel to the dedicated tbi table. 5x2: lines on skin to check abutting fields | MLC | we use a perspex spoiler screen covering the entire patient (200 x 50 cm, 2 cm thick). as close a possible to the patient. | none other than standard hospital hygiene | 2 dedicated technicians for setup, placement and readout of the mosfets and irradiation, and radiation oncologist on call for emergencies. | Pre-treatment checks are the standard checks as for any external beam treatment such as MU Check and checks bij technicians, physics and radiation oncologist (epid dosimetry is not possible because of field size). MuCheck. | none, our local experience. | Total body irradiation, toward optimal individual delivery: dose evaluation with metal oxide field effect transistors, thermoluminescence detectors, and a treatment planning system. van gurp et al 2007. ijrobp Non-myeloablative conditioning with allogeneic hematopoietic cell transplantation for the treatment of high-risk acute lymphoblastic leukemia Ron Ram, Rainer Storb, et al 2011 Haematologica | yes | SAFER | none
11 | a. paediatric and adult conditioning for allotransplant with stem cells or bone marrow for haematological conditions like ALL, AML, multiple myeloma, lymphoma b. 6x2Gy/3days, min. 6 hours interval; 2Gy/1day; 2x2Gy/1day, min. 6 hours interval; 8Gy/1day c. chemotherapy and/or ATG can be administered before, during or after the TBI, depending on the treatment protocol. In most cases transplant will be given 24 hours after irradiation. | No | On site | Regularly scheduled | 25 | ''Manual'' (combination of pdd tables and some kind of thickness determination) |  | 2 | 2 beam (anterior and posterior) number of MU &gt;2000 so three to four sub beams (identical but number of MU &lt;1000) | No |  | Gantry 86° patient is postioned perpendicular to the patient orientation | 15 | 17-20 cGy/min (6x2 Gy 8 hours interval; or 1x2Gy or 2x2Gy) 9-10 cGy/min (1x8Gy) | standing (AP-PA) with supporting aids such as saddle, armpit and hand supports. | focus-midline | 500 | 100 | 40 | 45 | MU determination using 0.6 cc farmer chamber in patiënt equivalent phantom. Verification of the transmission of custom made lung blocks (2%) | No |  |  |  |  | none, manual tables are used we only use density corrected path lungs for determination of the thickness of the lung shielding blocks | manual tables are used we only use density corrected path lungs for determination of the thickness of the lung shielding blocks | Yes | individual lung shielding based on planar imaging on film; eye shielding using lead blocks positioned according to beam light field if applicable | using euivalent path length correction and logarithmic absorption algorithm | double sided tape on plexi plates used for generating scattered low energy particles | point normalisation midline at the level of the umbilicus | no | standard plans and manual input of patient specific MU's | Standing supported at perineum, hand support, armpit support, head support, belts for head abdomen and pelvis | reference crosshair is mounted on fixed position on plexi scatter plates as reference for shielding (crosshair is visualised on film) | 4,8 m | no | gloves, mouth mask, protective clothing | 3 nurses (2 for positioning of patiënt, 1 operator of the machine ) (MD overal verification, positioning and shielding) (physicist, during simulation and first treatment, supervision) | technologist, making of the lung blocks physicist, MU determination using phantom measurements and lung block transmission measurements. Physicist, positioning of the blocks on the scatter plexis | AAPM17 | no | No
12 | Bone marrow transplantation 6x2Gy/3days 5x2Gy/3days 1x2Gy | No | No | Regularly scheduled | 10 | CT Based | Lateral left or right decubitus Start : upper part of the skull End : half of the legs | 2 |  | No |  | Gantry at 270, patient lying on their side on a special table, one anterior field, one posterior field, table is turned between the two fields | 18 | 6x2Gy with around 16cGy/min | Lying right side usually but some left side | focus-skin | 318 | 100 | 40x40 | 45 | No pre-treatment QA only invivo | Yes | Common diodes | centrally | 3% |  | Pencil beam algorithm/AAA (Eclipse) with heterogeneity correction | CT-based | Yes | Individual | 8cm | In front of the lungs on a support next to the table | Point (middle of the patient at the level of the ombilic) | No | Automatic : only the fields are duplicated manually to take into account the presence of shieldings : shielding is used half of the time for each beam | Like on the CT scan, legs slightly closed to enter the light field | Via CR digital images |  | Scatter perplex screen in front of the patient (1cm) |  | Nurses (=RTTs in Belgium) | Physicist | I have to check with the doctor | I have to check with the doctor | No
13 | leukemia, recurrent lymphoma and multiple myeloma only adults (paediatrics in Ghent) 6x2 Gy in 3 days, at least 6h interval, partial lung protection (75% of dose) (full allo transplant, younger pts) or 1x2 Gy (mini allo transplants) all TBI treatments are carefully scheduled according to haematological protocols (EBMT) in close cooperation with referring heamatologists | Yes | On site | Regularly scheduled | 20 | CT Based | sideways, total body ct, slice thickness 1 cm | 5 | 1*2Gy: 2 beams 6*2Gy: 5 to 6 beams | Yes | 1 mainly for dose reduction on the lungs | gantry at 90 degrees and the patients lying at their side, therefore AP/PA for the patient; patient is turned after 1 Gy | 10 | (600 MU/min) for all dose prescriptions | lying, left side | focus-midline | 400 | 100 | 40*40 | 45 | There is no daily pre-treatment QA specific for TBI. We consider standard QA sufficient. Verification of the TPS calculations on water phantom with measurements in water phantom with Farmer chamber. Measurements with and without MLC. Filmdosimetry and 2D array for FiF with MLC 3 %, 3 mm | Yes | diode | CAX | 10 % | yes, Sun Nuclear Isorad diode detector IBA DPD-3 electrometer | Model AAA 13.6.23 | CT-based, PerFraction | Yes | individual, MLC | MLC,field in field technique | MLC,field in field technique | point, isocentre, midpoint | Not standard, on individual basis. Only Dmax &lt; 130 % | automatic, Aria | vacuum matras, left side, knees bended, right arm stretched on the body, left arm folded around the head | -centre lightfield on marked isocentre patient -portable MV epid, Agfa CR25 | NA | spoiler-screen, 1 cm, full body: 200 cm * 100 cm, against the treatment table | desinfection of the hands with alcohol, alcogel | 2 - 3 RTTs: executing the treatment physician: judgement of the lung shielding physicist available for technical problems | only standard daily QA on Varian TB, no dedicated checks for TBI | AAPM 17 | Cem Onal M.D., Aydan Sonmez M.Sc., Gungor Arslan M.Sc., Serhat Sonmez M.Sc., Esma Efe M.Sc., Ezgi Oymak M.D.; Evaluation of Field-in-Field Technique for Total Body Irradiation; Int J Radiat Oncol Biol Phys 2012; 83: 1641 – 1648 | Yes | HFMEA general for radiotherapy | not performed for TBI
14 | Allogene (=donor) stemcelltransplants (multiple hematological malignacies): indications : AML, CML, MDS, Multiple myeloma, ALL, Folicular NHL, Mantelcelllymfoma, Aplastisch anemea. Myeloablativef: 3x4, 3x3.3 (future:6x2 Gy) (lung dose reduced) Non-myeloablative: 1 x 2 Gy (RIST=reduced intensity stamcel therapie) Stamcelltransplant within 24 hours. Interval between fractions is 24 hours (for 6x2 Gy: appr 8 hours). | Yes | On site | End of the day | 59 | CT Based | High dose - CT is used for lung block shaping and thickness; also to determine the patient thickness on different levels. Scan: 5mm coupes, crown till symphysis - PDD-tables (in Excel-sheets) are used for calculating MUs For 1x2Gy: No CT is used; only MU-tables in Excel-sheet for patient thickness |  | AP and PA. for high dose In AP: eye lens shielding In PA: lung shielding for 1x2 Gy: no shielding is used | No |  | high dose: gantry 90 (or 270); patient lying at their side, thus AP/PA for patient. 1x2 Gy: id, but patient is standing (or lying sideways if standing for 10 min is not possible) | 15 | All dose levels: approx 20 cGy/min midplane (10 cm depth) | high dose: lying, right side 1x2Gy: standing on block on the floor (or lying sideways on a table) | focus-midline | 500 | 100 | 40x40 | 45 | No pre-treatment QA. The standard linac QA-checks are sufficient. | No |  |  |  | no | NCS-report 12 + equivalent path length | Table (Excel) based using measured PDDs at 500 cm for different field sizes. The SPs are derived form SSD=100 cm. | Yes | Lungs: individual Eye lens: standard | Using a TPS-calculation (DVH's) of the lung dose without shielding. | Using an in-house system: movable positioning wheels on the TBI platform; There is s system using perspex / teflon shifting/rotation and the block is mounted on this system using a screw. | Point normalisation in the pelvic | The calulation spreadsheet gives an impression of the dose inhomogeneity. However no clinical decisions are done with it. The doses at the ankle could easily be &gt; 15 % more than the prescription dose. | Manually, individually in portions of 1000 MU. | High dose: left arms lying side ways and right arm above head (if possible) 1x2 Gy: arms side ways | high dose only: Lung shielding: done with Theraview TBI imager cf CT-reconstructed BEV with delineation of the shielding | high dose only: focus-shielding = approx 480 cm | high dose only: Yes: a perspex plate of 1 cm thickness as scatter-screen | no special precautions are needed | High dose 2 Radiotherapists: perform the radiation Radiation Oncologist: check poistion lung shielding on radiographic film Medical physicist: assist with lung and eye shielding 1x2Gy Radiotherapists: perform the radiation | High dose: Medical physicist: positioning and checking thickness of the lung blocks | AAPM Report 17 | NSC-12 | No |  |  | Number of patients per year: 2x6Gy: 4 (=20 in last 5 years) 1x2 Gy: 36 (=146 in last 4 years) Energy: 2x6Gy: 15 MV 1x2 Gy: 6 MV 1x2 Gy is scheduled regularly
15 | Indication: Fractionation scheme: 1 * 9 Gy (lung: 7 Gy) Time scheme: Criticality of timing: | Yes | On site | Start of the day | 1 | ''Manual'' (combination of pdd tables and some kind of thickness determination) |  | 2 |  | No | 1 To reduce lungdosis from 9 Gy to 7 Gy lungshielding is applied in both beams and eyeshielding is applied as the patient is facing towards the beam. | Gantry 90 degrees, Collimator 45 degrees, maximum fieldsize. Patient at right and left side position. Legs bended and arm/hand wrapped around the head. | 6 | +- 37 cGy/min midplane (600 cGy/min * (101,6/368,6)^2 * 0,8175) | Patient at right and left side position. Legs bended and head bending towards chest and arm/hand wrapped around the head. | focus-midline | 377 | 100 | 40 * 40 | 45 | routine dosimetry within 1 % | No |  |  |  | no | ''Manual'' (combination of pdd tables and some kind of thickness determination) For lung thickness correction an average density of 0,25 g/cm^3 is assumed. | Manual MU calculation. Detph doses and outputfactors have been measered at TBI distance. The standard condition is a total thickness of 20 cm and a midplane thickness of 10 cm. For the number of MU's we apply a correction factor to correct for the real thickness of the patient at the level of navel. | Yes | individual shielding of the lungs and a standard shielding for the eyes | We detemine the effective tickness of the body at lung level by measuring, on CT images, the total thickness of the thorax minus 0,75 * lung thickness. This effective thickness of the thorax is used to calculate the midplane dosis. And then we determine the thickness of the shielding by searching in a look-up table (thickness versus transmission) the thickness of the shielding which gives the desired transmission to achieve 7 Gy midplane dosis. | The lung shielding is taped on plastic plates | The midplane point at te level of navel is our description point. | For a number of planes midplane dose is calculated. In case the midplane dosis of the head is to high the intended midplane dose of 9 Gy may be lowered. Tolerances?? | Manual | Patient at right and left side position. Legs bended and head bending towards chest and one arm/hand wrapped around the head and the other along the body. The belly is | We use an diagnostic phosphor imaging plate, which is placed behind to patient to verify the placement of the lungshielding. | The lung shielding blocks are placed at a distance of 2 m from focus. | no | none | 1 radiation oncologist, 3 RTT's and a medical physicist. | Radiation oncologist: approve dose calculation RTT's: checking (immobilisation) devices, Medical Physicist: checking MU calculation | none, as we started (erly 90's) we did some site visits and that was it and still is. | we use the same protocol as we started (nothing has been changed). Publications, a lot, has up till now never led to a change in our protocol. But there was a publication of TBI with VMAT. Maybe this is something for the nearby future. | No
16 | 'Before BSCT (blood stem cell transplantation, autologous or allogenic) in the treatment of Leukaemia (AML) or Kahler's disease or lymphoma. A fractionation scheme of 6 x 2 Gy (three consecutive days, twice daily), 3x3,3 Gy (on consecutive days), 2 x 4,5 Gy (on consecutive days) or 1 x 2 Gy. Occasionally, a 2 x 2 Gy schedule is used (on a single day, minimal 6h interval). chemotherapy is given before radiation treatment, transplantation is performed the day after TBI. | No, however for the new technique we still have a dedicated team but the aim is to train all RTTs | On site | Scheduled ad-hoc | 40 | CT-based for high dose and manual for low dose (2/4 Gy). The aim is to change to CT-based for all patients. | New technique: Patient is positioned on right side with knees bent (max. scan length 160cm). Slice thickness 5mm | 14-16 | New technique: A/P forward IMRT technique; 7-8 beams (with control points) for both A/P, Shielding of lungs is done with MLC | Yes | Both lungs are shielded from both A/P with MLC leaves; additional beams to homogenise dose | Gantry at 273,4 degrees (reason for not using 270 degrees; TBI table does not reach isocenter height and patients is positioned 15cm lower where additional lasers are placed); patients lying on right side; AP/PA beams | 10 | 33 cGy/min in parts unblocked by MLC leaves | lying on right side on dedicated table | focus-midline. Midline is defined as the middle of the TBI table | 350 | 100 | 40 | 45 (for open field, for IMRT segments angle can be different) | Independent secondary dose calculation with Mobius3D. Point dose should be within 5%; 3D gamma calculation cannot be used due to bug with contamination electrons at high-SSD in Mobius3D | Yes | Mosfets | Currently (implementation phase): neck, lower back, level of left lung, ventrally and dorsally; We aim to decrease the number of measurements in the future | 10 % | No | Pinnacle Adaptive Convolve (collapsed-cone) with inhomogeneity correction | CT-Based | Yes | individual using MLC segments | TPS (field in field block with MLC) | NA (MLC/Gantry) | DVH, 3D dose distribution | Yes: For low dose: 90-130%, for high dose 80-120% | automatic | Patient is positioned lying on right side in vacuum mattress, legs bent, left arm lying on side, right arm bent with hand on top/back of head. Lines drawn on skin during CT session to assist in positioning. In TPS: beam is created that follows patient contour. This is used for a first check of positioning using the light field (after positioning the patient using the lasers and lines on skin). This beam is not irradiated; | Fields are set up and irradiated covering the lungs with a small margin to check positioning with respect to the lungs using a Cablon mobile EPID panel. | >300cm (MLC is used) | a spoiler screen (10mm perspex) is placed as close to patient as possible. | 'personnel wears special clothing and masks. Treatment room and all parts (table/MOSFET dosimeters) are cleaned before treatment. | During implementation phase both a physicist and radiation oncologist are present during treatment for interpretation of EPID images and MOSFET measurements. Three RTTs are present. | Physicist and RTT check treatment plan (planning technique and dose distribution). Radiation Oncologist checks dose distribution. | none | amongst others Bär et al, Bone Marrow Transpl. 10(1) p45, 1992 | Yes | HFMEA
17 | o 6*2Gy (12 Gy mid-abdomen, 8 Gy mid-lungs) in 6 fractions (Tuesday PM, Wednesday AM and PM, Thursday AM and PM, Friday AM with min. 6 hours delay between AM and PM treatments) o 1*2 Gy | No | Other site | Regularly scheduled | 23 | a standard VMAT technique (1ARC ant and 1 ARC post per fraction with lung Pb blocks, patient lying on a treatment couch on the floor, UM depending on mid-abdomen thickness) |  | 4 | Until end 2015 : Ant and Post (lateral position) for 6 fractions. 1 Ant and 1 Post field for the 2 last fractions with Pb blocks to protect the lungs. From January 2016 : Ant and Post Arcs | Yes | 1 Until end 2015 : Ant and Post fractions with Pb blocks are given in 3 contiguous sections. | Until end 2015 : gantry 90°, patient lying on their side for 6 fractions. Gantry 0° for the 2 last fractions with Pb blocks, patient in supine and prone position. From January 2016 : 130° Arc in Ant and Post position (supine and prone position) | 6 | Until 2015 : 1x8Gy (midplane) in min. 6 hours From January 2016 : 6x2Gy (midplane) in 3 days | Until 2015 : left side or right side for 6 fractions. Supine and prone position for the last 2 fractions From January 2016 : Supine and prone position for each fraction | focus-midline | about 2 m (sweeping) |  | 40x40 (max. field size) | Until 2015 : 45°, From January 2016 : 0° | Dose measurement must be max.2% away from expected dose. | Yes | TLD | on abdomen and spoiler |  | In-vivo dosimetry is used for each patient | NA | MU are calculated based on a measued depth dose curve (depending on the patient thickness) | Yes | individual Pb blocks for lungs | fixed thickness based on measurements | Until 2015 : on the spoiler From January 2016 : on the patient | NA | no | Until 2015 : no R/V system used From January 2016 : Automatic and manual method | Until end 2015 : patient lying on his side with mattresses between 2 spoilers, 1 arm on his side and the other in front. For the 2 last fractions, the patient is lying supine or prone, his arms along his body From January 2016 : patient lying supine or prone on a small mattress, his arms along his body. A spoiler on top of the bed. | X-Ray | Until end 2015 : the Pb blocks are positioned on the spoiler. From January 2016 : Pb blocks are positioned on the patient | Until end 2015 : the patient is lying between 2 plexi spoilers (4cm thick) From January 2016 : a 1 cm plexi spoiler plate is positioned 40cm on top of the patient | Until end 2015 : depending on patient (everything is cleaned in advance, use of masks and gloves) From January 2016 : everything to be cleaned | - 1 radiotherapist (check of dose and positioning of patient and Pb blocks) - 2 RTT (in charge of patient and dose delivery) - 1 physicist (positioning of patient and dose monitoring) | - Radiotherapist : CT scan, measurement of thickness of patient, contouring lungs on CT scan - Dosimetrist : manufacturing of Pb blocks and preparation of treatment calendar in Mosaiq - Physicist : UM calculation, dosimetry before treatment (chamber and diodes) | Until end 2015 : ? technique used since 1985 ? | From January 2016 : ''Arc therapy for total body irradiation-A robust novel treatment technique for standard treatment rooms'', A. Jahnke., Radiotherapy and Oncology 110 (2014) 553-557 and ''Accelerating total body irradiation with large field modulated arc therapy in standard treatment rooms without additional equipment'' M. Polednik, Strahlenther Onkol (2015) 191:869-874 | Yes |  | From January 216 : failure in arc delivery
TSI
inst nr | Question 55 : For which indications is TSI used in your institute? | Question 56 : Do you have a dedicated TSI team? | Question 57 : Do you have a backup treatment machine | Question 58 : How many patients do you treat annually with TSI? | Question 59 : Do you have to allocate a time slot for treatment? | Question 60 : Briefly indicate method of planning (more elaborate in section further on) | Question 61 : Number of large fields | Question 62 : If necessary: More information on the number of large beams | Question 63 : Additional fields (e.g. crown, shoulder, foot, perineum) and relative weight/dose) | Question 64 : Specify beam directions and time pattern (e.g. given on alternate days) | Question 65 : Beam quality/energy (Mev). Multiple possible? | Question 66 : Dose rate on patient skin (cGy/min). Differentiate per dose level if necessary (eg 30 Gy, 36 Gy). | Question 67 : How is the patient positioned (e.g. standing, lying,..) | Question 68 : Field sizes of large and additional fields | Question 69 : Field orientation (collimator angle) | Question 70 : What organs are shielded and how is this realised? | Question 71 : Describe your Pre-treatment QA: what tolerances are allowed and what actions are performed. | Question 72 : Is in vivo dosimetry performed on a routine basis? | Question 73 : What detectors are used for in vivo dosimetry? | Question 74 : What positions are measured in vivo? | Question 75 : What are the tolerances on the in vivo measurements? | Question 76 : Do you use In-Vivo dosimetry non-routinely (eg. while introducing a new technique)? Please elaborate. | Question 77 : Which calculation algorithm is used (or manual or table-based calculation or MU calculation) please elaborate. | Question 78 : How is the planned skin dose verified? | Question 79 : How is the dose distribution normalised (point)? | Question 80 : What are the tolerances on the dose inhomogeneity? | Question 81 : How is the treatment data entered into the R/V system (automatic, manual, combined (both manual and automatic), standard plans) | Question 82 : How is the patient setup verified? (including positioning of extremeties/hands) | Question 83 : How is the position verification of shielding and other tools/devices performed? | Question 84 : Describe distances between machine and shielding devices | Question 85 : Do you use additional beam modifying devices (e.g. “reflector/scatter plate” above head)? Please specify construction (sizes, materials used, position) | Question 86 : What published guidelines did you use (AAPM, IAEA, etc..) | Question 87 : Have you used scientific publications (please state which)? | Question 88 : Have you performed a prospective risk analysis while setting up your treatment? (include e.g. lasers, cross-hairs,..) | Question 89 : Which method of risk analysis did you use? (e.g. HFMEA, ..) | Question 90 : Which prominent risks were indicated in the analysis? | Question 91 : Please add here any additional relevant information.
1 | a. Cutanuous lymphoma 12Gy at 2.1mm depth b. 8x1.5Gy, 1x/day c. Does not apply | Yes | On Site | 6 | Regularly scheduled | All patients receive the same plan-design. Dose prescription is made at 2.1mm depth and mainly 8x 1.5Gy. | 2 | 2 fields of 9MeV, size of 34x34cm² at iso with G252° and G288° are delivered for each one of the 6 orientations of the patient (as described in AAPM report 23, standford technique). | Additional fields are determined by clinical visible recurrence after the treatment with large fields. The prescription takes into account the in-vivo dosimetry results from the treatment with large fields. | Beams for all 6 patient-orientations are delivered each day. | An energy degrader of 2cm plexiglas is placed 20cm before the surface of the patient, using a 9MeV electron beam. Mean entrance energy: 3.15MeV R50: 13.5mm | 24.9cGy/min. Each part of the patient is exposed by 3 of the 6 orientations. In order to decrease the treatment-time as much as possible, the dose-rate of machine is set on 1000MU/min. | Patient is standing and take the 6 orientations described by the AAPM-report-23 ''six dual-field beams''. Rotation of the patient is splitted onto 6 orientations, each one implies a rotation of 60° from the previous one (6x60°=360°). Patient is positioned at 3,48m from the source. | Dual fields of 34x34cm² at iso with G252° and G288° are used to generate an homogeneous field (&lt;8.5% surface variation) of 1m75 height and 50cm width. | 0° | There is no shielding. | There is no specific pre-treatment QA for TSI. | Yes | Gafchromic EBT3 film | 16 pieces of 5x2cm² of Gafchromic EBT3 films are positioned on relevant anatomic regions: head, CAX, perineum, top of feet, sole of feet, hands, shoulders, neck, arms, legs, etc. | 10%. Anatomic regions are expected to be out of tolerances (sole of feet, perineum, top of head, etc.) and the in-vivo measurement is taken into account for ulterior boost. | Gafchromic EBT3 film are used for new techniques (surface dosimetry). | / | The planned skin dose is determined by measurement during commissioning of the technique with IC and film dosimetry. In-vivo dosimetry is performed during the first fractions for each patient. | At 2.1mm depth at beam center of a horizontal beam projection (gantry 90°). | 10% for surface dose (as recommended by the AAPM report 23). | Manually performed from a standard plan | The patient setup is visually controlled by the RTTs. | There is not shielding and the scatterer/energy degrader is visually controlled by the RTTs (marks on the ground). | Scatterer/energy degrader is positioned at 3m from the exit window of the gantry. | The scatterer/energy degrader is actually the 2cm plexiglas plate used for TBI. The plate is placed 20cm before the skin of the patient. | AAPM report 23 | / | No |  |  | We controlled ozone production during the irradiation. Intensive irradiation (12048MU on 30') did not lead to noticeable ozone concentration level.
2 | mycosis fungoides: 20x1.75 Gy, 5x2 Gy for reirradiation, 6x2 Gy for low dose TSI | Yes | On Site | 15 | Regularly scheduled | MU table and standardized fields | 12 | we apply 6 treatment positions and two gantry angles (upper and lower half body). If &lt;=10 fractions all 6 treatment positions are applied each fraction. If &gt;10 fractions are given only 3 of 6 positions are given each day. | crown, shoulder, foot, perineum, sometimes mamma. All additional fields get 50% of the total dose | modified Stanford technique | Always 4 MeV | HDRE produces a 10 times higher doserate. The patient is at 3m hence approximately 330 cGy/min. (note total procedure takes longer since only part of the patient is in one fiel, 12 large fields plus additional fields are required to cover the total skin) | standing | midpatient at 3 m hence large fields 120 cm x 120 cm | 0 degrees, fields need to be matched | eye lense shielding using lead goggles | no pretreatment QA, only standard linac QA | No | 0 |  | 0 | yes we have performed film measurments on patient and rando phantom | manual table based MU calculation | standardized dose, class solution verified using phantom measurments | normalized to the average dose in the large fields | ca. +/- 15% | plan setup on water phantom in Pinnacle and from Pinnacle sent to Mosaiq | patient stands in a frame and is visually checked by radiation therapists | we only somtetimes use lead goggles and or lead shoes, nno further position verification on this | the shielding is on the patient, hence approximately 2.9 m from focus since the mid patient is at 3 m | no we do not | AAPM report 23 |  | No
3 | a. mycosis fungoides, exceptionally other skin lymphomas b. 7-10x3Gy/every other day c. n.a. | Yes | On Site | 3 | Regularly scheduled | no planning, commisioning measurements on phantom serve as basis and adaptation using in vivo measurements during first session | 12 | 12 beams, abutting on cranio-caudal direction, overlapping left right without applicator | perineum, armpits, groins and if necessary crown and feet | every other day -25, +25° degrees on a stretcher, patient in prone and supine position, 3 abutting fields in cranio-caudal direction so 12 fields in total | 4 MeV | 68 cGy/min | lying on a stretcher | 40x40 cm² at SAD, but at 170 cm SSD (so 68x68cm²) boost fields most often 10x10cm² at SSD=95 cm | 0 for large fields, invidual setting for boost fields | eyes (lead shielding on eye ball) | measurement on phantom (patient eguivalent size) at dmax for MU determination | Yes | radiochromic film for skin dose, diodes for dose at dmax | head, arms, thorax, adbomen, legs, feet, armpits, crown, groins, perineum, back between the shoulders, lower back region | 10 % | identical as on patient level | no, manual calculation and verification by measurements on flat chambers | radiochromic films | at 4 mm depth | 20 % for individual points (average spread &lt; 10%) | manual using standard plans | supervision by MD (every treatment) and physicist (first treatment) | supervision by MD (every treatment) and physicist (first treatment) | shielding devices are at skin level | no, only lead shielding for excluding skin regions not to be exposed | modified protocol starting from AAPM 23 | no | No
4 | Mucosis fungoides x fractions of 2 Gy + ''bandelets'' (small fields on hidden parts of the skin) | No | No | 2 | Regularly scheduled | 0 | 12 |  | If needed | 45 deg | 6MeV+perpex diffuser | around 400 | Lying | 40x40 | 0 | Eyes with lead lenses | No | No | 0 |  | 0 |  | Same MU for all patients | Same MU for all patients | Same MU for all patients |  | Manually | Light field + ground marks | Light field | On the patient | Diffusing plate of perpex | I have to check with the doctor | I have to check with the doctor | No